F20 Quick Checklist

Start up:

- Cool down the microscope

- Center the stage before inserting specimen holder

- Once the specimen holder is inserted, do not open the column valves until the IGP reading is below 20

- Turn on ‘Low Dose’ function

- In low dose ‘Search’ mode, check the z height

- Switch to ‘Focus’ mode and press the eucentric height button

- Switch to ‘Exposure’ mode

- Check condenser and objective aperture alignments

- Under ‘Direct Alignments’ check gun tilt, gun shift, pivot point x and y, beam shift and rotation center alignments

- Switch to ‘Search’ mode and align ‘Search’ and ‘Exposure’ modes

- Collect new gain reference for K2 in Digital Micrograph using beam settings for ‘Exposure’ mode (see [Instructions for Gatan K2 Camera])

- Adjust K2 configuration in Digital Micrograph for using either UCSF Image4 or SerialEM

- Set up 8 bit image stack collection for USCF Image4, if necessary

- Begin imaging

Shut down:

- Close the column valves

- Center the stage and remove specimen holder

- Log out of your account and log into guest account

  - Username: F20guest

  - Password: walzlab

- Start the ‘Cryo Cycle’, remove the liquid nitrogen dewar, place Styrofoam box beneath the coils and safely store the dewar on the desktop

- Transfer all of your images from the GPU to your own account/computer

- Heat the K2 to 50°C, set the time to as much as permitted. Leave one hour for k2 cooling down to -20°C before next user start.

- Sign the log book and note any problems

If you encounter any problems, or think anything is strange with the microscope, please contact Zongli or Melissa before continuing.